Visual Threshold Testing Requirements

Jenks Vestibular Physiology Lab

June 25, 2025

Minimum criteria for success are highlighted in green.

To be determined: The participant will be X cm away from the M x N cm displayed image. (This will depend on if the image is displayed with the projector or a laptop). The field of view should be >40 deg (likely 60 deg), based on the findings in Karmali 2014.

| 1 | Physical Hardware | The test will involve an image (to be provided) displayed with high resolution on a projector or laptop screen. The participant will be seated in the (unmoving) NKI chair with their head secured and the lights off. A draping over their head will connect to a round (for tilt) or square aperture (for translation) placed between their face and the image to remove any visual stimulus from the surroundings. (Rachel and Faisal will develop the round and square aperture hardware). |
| --- | --- | --- |
| 2 | Directions of motion | Four modes should be included in the software. The image will move up or down (Z/vertical), left or right (Y/horizontal), rotated left or right (roll), or zoomed in/out (X/fore-aft), based on an axis passing through the cyclopean eye. |
| 3 | Display | On the screen/projection, the image will be displayed as a circle or square (for roll tilt) or a rectangle (for translations). (Rachel and Faisal will provide the cropped images). The image will be bordered by black pixels on the screen/projection (the image should never reach the edge of the screen during movement stimuli). The image will rotate on the screen or translate across the screen. Importantly, we are not changing the windowing of the image, we are moving the entire image (simulating a scenario where the human is moving, and the image is staying still). |
| 3 | Single-cycle sinusoid of acceleration | The movement of the image will be a single-cycle sinusoid of acceleration (see MOOG and NKI chair threshold testing programs). After each stimulus, the image will hold the final position. Movement amplitudes will be specified in physical units (degrees, cm, deg/s, cm/s) and the appropriate geometry should be applied to get from physical units to pixels (e.g., distance to display) |
| 2 | Randomized direction of movement | Each stimulus will move in one direction. Direction will be randomized on every repetition (see MOOG and NKI chair threshold testing programs). |
| 3 | Adjustable start amplitude | The software should include an adjustable start amplitude (i.e., the maximum position of the first stimulus). The default should be 2 degrees or X (TBD) cm. |
| 3 | Adjustable frequency | The software should include an adjustable stimulus frequency (i.e., the frequency of the acceleration sinusoid). The default should be 1 Hz, with an allowed frequencies range from 0.05 to 5 Hz. |
| 4 | White noise | During the stimulus, white noise audio will play from the computer speaker. |
| 5 | Responses from buttons | Participants will provide a response (i.e., indicating which direction they perceived movement) using two push buttons (push buttons, computer keyboard, or mouse buttons). |
| 6 | Stimulus amplitude adjusted in real-time using PEST staircase | The software will adjust the amplitude of the next stimulus in real-time, using the participant responses and PEST staircase. |
| 7 | Adjustable PEST staircase settings | The software should include adjustable PEST staircase settings (i.e., m and n should be adjustable for m-down and n-up). The default should be 4-down and 1-up (see MOOG and NKI chair threshold testing programs). |
| 4 | Pause duration | After end of the stimulus, the image will be held in its final position for X seconds. The software should include an adjustable pause duration. The default should be 2 seconds. |
| 6 | Adjustable number of repetitions | The software should include an adjustable number of stimuli repetitions. The default should be 100 repetitions. |
| 7 | Start, Stop, & Pause | The software should include a button that allows the operator to start testing. There should also be a button that allows the operator to pause or stop testing. If pause or stop are selected during a stimulus movement, the pause or stop will occur after the completion of the ongoing stimulus. After pressing the pause button, the operator should be able to resume testing. |
| 8 | Participant Code & Condition | Before pressing start, the operator should be prompted to enter a participant code (to be used for file naming). There should also be an optional field to input a condition description (e.g., “XTrans1Hz”). |
| 9 | Display user responses vs actual magnitude | The software should display a running list of each repetition, showing the participant response (+1 or -1) and the actual plate position (+/-X degrees/cm) |
| 10 | Reaching edges of the image | While the image does not need to return to a “home” position after each stimulus, we will need a solution for if the stimulus approaches amplitudes near the limits of the screen. |
| 12 | Save data | The participant responses and actual image positions should be exported into a file named “yyyymmdd_hhmmss_ participantcode_condition”. Where the time is the recorded when the operator presses start. Data should be exported automatically (even if the “Stop” button is pressed). The operator should be able to adjust the save location prior to starting the test. |